## USB Packets
USB Packets
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64
Hex | 00 | 00 | 00 | 3c | Payload
1-3 | Prefix, at the moment always 0x000000 for outgoing, or 0x414243 ("ABC") for incoming
4 | USB Frame length in bytes, prefix and itself excluded. Unsigned 8 bit integer
5 - 64 | USB Message payload. Payload is always 60 bytes. If Byte 4 indicates anything smaller than 0x3c, then the remainder of the payload is null padded.
If Byte 4 indicates a size of 0x3c, it's possible that the next (and further subsequent) packet(s) to be received will be concatenated together with this one to create one full Bayer Message
The payload forms the message body of the other tabs in this spreadsheet

## Bayer Command
Bayer Command
Contained by: Single USB Packet
Byte | 1
Hex | 58
1 | Bayer Command.
0x58 ("X") - Query from computer
0x05 - (ASCII ENQ) - Enquiry
0x15 - (ASCII NAK) - NAK
0x04 - (ASCII EOT) - EOT
0x06 - (ASCII ACK) - ACK
The EOT from the Contour may also have a valid ASTM packet sequence that follows it. Eg, 0x040x020x23…
See the last page of https://github.com/bewest/diabetes/blob/master/bayer/MeterInterface.pdf for a more detailed description on entering and leaving command mode

## Bayer Capability Query
Bayer Capability Query
Contained by: Single USB Packet
Byte | 1 | 2
Hex | 57 | 7c
1 | Capability Query. Either 0x57 ("W" - assume write commands), 0x51 ("Q" - assume query commands) or 0x31 ("1" - assume magic commands)
2 | Command delimiter. 0x7c ("¦")

## Bayer Binary Message
Bayer Binary Message
Contained by: One or more USB Packets in 60 byte chunks
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | .. | N
Hex | 51 | 03 | 30 | 30 | 30 | 30 | 30 | 30 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 12 | 04 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 15 | 00 | 00 | 00 | fe | Payload
1 | Command Action, 0x51 ("Q") Query
2 | Command delimiter. 0x03
3-8 | Pump serial number. For pumps older than the 640G, this returns a 6 digit ASCII serial number. For the 640G, it returns "000000"
9-18 | Unknown bytes, currently assume always 0x00
19 | FrameType
0x10 - OPEN_CONNECTION
0x11 - CLOSE_CONNECTION
0x12 - SEND_MESSAGE
0x14 - READ_INFO
0x16 - REQUEST_LINK_KEY
0x17 - SEND_LINK_KEY
0x80 - RECEIVE_MESSAGE
0x81 - SEND_MESSAGE Response
0x86 - REQUEST_LINK_KEY Response
20-23 | Frame type sequence number. Currently assuming a little-endian 4 byte unsigned int
24-28 | Unknown bytes, currently assume always 0x00 padded
29-32 | Payload length in bytes and over frames. Length >27 means that payload continues in next incoming frame (1-n). Currently assuming a little-endian 4 byte unsigned int
33 | Binary Message CRC: It’s the lower byte of the rolling checksum of the entire packet from the Q up to the end of the payload but NOT including that 1 byte itself
34-N | Payload of the Bayer Binary Message. Length is determined by the Payload Length defined in bytes 29-32

## Medtronic Message
Medtronic Message
Contained by: One or more Bayer Binary Messages
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | .. | N | +1 | +2
Hex | 55 | 1c | Payload | [53 empty] | CRC
1 | Command Action
0x03 - Pump comms channel negotiate
0x05 - Pump request
0x55 - ("U") - Pump response
2 | Medtronic message size. Size of the message body, from byte 1 until the end of the payload, but not including the CRC
3-N | Payload of the Medtronic Message. Length is determined by the Payload Length defined in byte 2. Though theoretical max is 255, max seen so far is 124 (7c)
N = value of byte 2
N+1 | CRC Least Significant Byte
N+2 | CRC Most Significant Byte
CRC is CRC-CCITT (0xFFFF)
In the Medtronic message, the CRC-CCITT is encoded as a little-endian 16-bit unsigned integer.
See lammertbies.nl/comm/info/crc-calculation.html

## OPEN_CONNECTION Request
OPEN_CONNECTION Request
Contained by: One Bayer Binary Message (Frame Type 0x10, from computer to USB stick)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
Hex | Payload
1-32 | Message Payload. Is a 32-byte HMAC - exact mechanism currently unknown, but can be retrieved from carelink.minimed.eu/com with a login
After logging in and storing the cookies, the request to carelink.minimed.eu goes to https://carelink.minimed.eu/patient/main/../secure/SnapshotServer/
The request should be a POST with a Content-Type of application/octet-stream
The POST body should be a Java ObjectOutputStream with a serialised int of value 0x1c, followed by a serialised string of the stick model and serial number, as returned by the Bayer ASTM message of format "6213-XXXXXX"
The response is a Java ObjectInputStream which contains a serialised array which is the HMAC to be used (after reversing the order of the byte array)

## OPEN_CONNECTION Response
OPEN_CONNECTION Response
Contained by:
One Medtronic Message (Command Action 0x55)
Which is contained by One Bayer Binary Message (Frame Type 0x10, from USB stick to computer)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42
Hex | 00 | 00 | Pump SN |  |  | ee | 45 | 02 | Pump MAC | [7 empty] | 82 | 00 | 00 | 00 | 00 | 00 | 07 | 00 | 5d | Link MAC | [7 empty] | 42 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 0e
1-2 | Null padded. May have some intrinsic meaning to the Medtronic message
3-5 | Pump serial number as a little-endian 3 byte integer
6-8 | Unknown bytes
9-16 | Pump MAC as a little-endian long. Pump MAC = Pump Serial + 0x0023F745EE000000
17-25 | Unknown bytes
26-33 | Link (USB stick) MAC as a little-endian long. Link MAC = Link Serial + 0x0023F70682000000
34-42 | Unknown bytes

## READ_INFO Response
READ_INFO Response
Contained by: One Bayer Binary Message (Frame Type 0x14, from USB stick to computer)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19
Hex | Link MAC | [7 empty] | Pump MAC | [7 empty] | 01 | 00 | 01
Outer Message
1-8 | Link MAC as a big-endian long
9-16 | Pump MAC as a big-endian long
17-19 | Unknown bytes

## Negotiate Channel
Negotiate Channel
Contained by:
One Medtronic Message (Command Action 0x03)
Which is contained by One Bayer Binary Message (Frame Type 0x12, from computer to USB stick)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26
Hex | 01 | 14 | 00 | 00 | 00 | 07 | 07 | 00 | 00 | 02 | Link MAC | [7 empty] | Pump MAC
1 | Sequence Number (always 1 for this message, even if there are multiple negotiation attempts)
2 | IEEE 802.15.4 Channel ID
0x0e - Channel 14 - 2420MHz
0x11 - Channel 17 - 2435MHz
0x14 - Channel 20 - 2450MHz
0x17 - Channel 23 - 2465MHz
0x1a - Channel 26 - 2480MHz
3-10 | Unknown bytes
11-18 | Link MAC as a little-endian long
19-26 | Pump MAC as a little-endian long

## SEND_MESSAGE Request
SEND_MESSAGE Request
Contained by:
One Medtronic Message (Command Action 0x05)
Which is contained by One Bayer Binary Message (Frame Type 0x12, from computer to USB stick)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | .. | N
Hex | Pump MAC | [7 empty] | 05 | 11 | 09 | Encrypted Payload
1-8 | Pump MAC as a little-endian long
9 | Sequence number
10 | Unknown byte
11 | Encypted Payload size - unsigned 1 byte int
12-N | Payload encrypted in AES/CFB/NoPadding. Size of payload is the value of byte 11
Exact key generation mechanism yet to be determined, but can be retrieved from carelink.minimed.eu/com with a login
After logging in and storing the cookies, the request to carelink.minimed.eu goes to https://carelink.minimed.eu/patient/main/../secure/SnapshotServer/
The request should be a POST with a Content-Type of application/octet-stream
The POST body should be a Java ObjectOutputStream with a serialised int of value 0x1c, followed by a serialised string of the stick serial number of format "XXXXXX"
The response is a Java ObjectInputStream with 3 serialised pieces
The first is an int (which must be more than zero - exact reason unknown, but that's what the applet code says)
The second piece is a serialised byte array with the encryption/decyption key
The third piece is an int, which means "count" according to the Java applet code, but the count of what is currently unknown
The IV is the same bytes as the Key, but with the first byte changed to the negotiated radio channel

## SEND_MESSAGE Response
SEND_MESSAGE Response
Contained by:
One Medtronic Message (Command Action 0x55)
Which is contained by One Bayer Binary Message (Frame Type 0x81, from USB stick to computer)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
Hex | 00 | 04 | 00 | 00 | 00 | 00 | 03 | 00 | 01 | 05 | 02
NB: Medtronic Message payload is empty if it's a response to a "Negotiate Channel"
1-2 | Command response code?
3-9 | Unknown bytes
10 | Sequence number. Should be the same as the corresponding SEND_MESSAGE (0x12) sequence number
11 | Unknown byte

## RECEIVE_MESSAGE
RECEIVE_MESSAGE
Contained by:
One Medtronic Message (Command Action 0x55)
Which is contained by One Bayer Binary Message (Frame Type 0x80, from USB stick to computer)
Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | .. | N
Hex | 00 | 06 | Pump MAC | [7 empty] | Link MAC | [7 empty] | 5a | 46 | 04 | 0e | Encrypted Payload
1-2 | Command response code?
3-5 | Pump MAC as a little-endian long
11-13 | Link MAC as a little-endian long
19-21 | Sequence Number as a little-endian 3 byte integer? 3 byte int seems unlikely, but that's best guess at this stage
22 | Encypted Payload size - unsigned 1 byte int
23-N | Payload encrypted in AES/CFB/NoPadding. Size of payload is the value of byte 22
Key generation mechanism information is available in the "SEND_MESSAGE Request" tab
The IV is the same bytes as the Key, but with the first byte changed to the negotiated radio channel